## Sheet1
| Sequence steps |  | Volumes/flowrates | Notes
Activate instrument | Turn on analyzer
Turn light off
Get dark scan (see seabreeze code)
Turn lamp on
Wait 3 minutes for lamp to warm up
Run Blank (2x) | Get reference scan (sea seabreeze code)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 600,60,1 | Dispensing blank sample (carrier)
COV-6 go to port 3 (Molybdate reagent)
| Pump 2 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 400,40,0 | Pumps run simulatenously
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 320,32,1
Wait 2 seconds
COV-6 go to port 5 (Ascorbic acid reagent)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 400,40,0 | Pumps run simulatenously
| Pump 2 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 320,32,1
Wait 2 seconds
Go to port 2 (flow cell)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 400,25,1 | ***depends on flow cell length (used 300 with 10cm FC)
| Wait 5 minutes |  |  | **Could be shortened given no PO4 in blank
Get absorbance value of blank sample (see seabreeze code)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 1000,150,0 | Pumps run simultaneouslty to flush instrument
| Pump 2 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 1000,150,1
Run PO4 standards (2x) | Get reference scan (sea seabreeze code)
COV-6 go to port 6 (PO4 standard)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 600,60,1
COV-6 go to port 3 (Molybdate reagent)
| Pump 2 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 400,40,0 | Pumps run simulatenously
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 320,32,1
Wait 2 seconds
COV-6 go to port 5 (Ascorbic acid reagent)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 400,40,0 | Pumps run simulatenously
| Pump 2 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 320,32,1
Wait 2 seconds
Go to port 2 (flow cell)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 400,25,1 | ***depends on flow cell length (used 300 with 10cm FC)
Wait 5 minutes
Get absorbance value of PO4 standard (see seabreeze code)
| COV-6 go to port 3 (Molybdate reagent) |  |  | Priming step: added 4/4/22
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 150,50,1
COV-6 go to port 6 (PO4 standard)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 150,50,1
Go to port 2 (flow cell)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 1000,150,0 | Pumps run simultaneouslty to flush instrument
| Pump 2 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 1000,150,1
Run PO4 seawater sample (1x) | Get reference scan (sea seabreeze code)
Run auxiliary pump for 60 seconds
COV-6 go to port 4 (PO4 sample)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 600,60,1
COV-6 go to port 3 (Molybdate reagent)
| Pump 2 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 400,40,0 | Pumps run simulatenously
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 320,32,1
Wait 2 seconds
COV-6 go to port 5 (Ascorbic acid reagent)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 400,40,0 | Pumps run simulatenously
| Pump 2 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 320,32,1
Wait 2 seconds
Go to port 2 (flow cell)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 400,25,1
Wait 5 minutes
Get absorbance value of PO4 sample (see seabreeze code)
| COV-6 go to port 3 (Molybdate reagent) |  |  | Priming step: added 4/4/22
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 150,50,1
COV-6 go to port 4 (PO4 sample)
| Pump 1 aspirate | volume (µL), flow rate (µL/s), hold all (0/1) | 150,50,1
Go to port 2 (flow cell)
| Pump 1 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 1000,150,0 | Pumps run simultaneouslty to flush instrument
| Pump 2 dispense | volume (µL), flow rate (µL/s), hold all (0/1) | 1000,150,1
post-processing | Calculate PO4 concentration:
[PO4] = (Abs_sample - mean_Abs_blank) x [concentration of po4 standard] / (mean_Abs_standard-mean_Abs_blank)
Store sample time stamp + sample concentration, send to server
Sleep until next sample (~40 min if wanting hourly measurements)